
氢化物发生器HS50操作规程（以汞为例）
注：测定不需要点火，在常温下进行测定
一、试剂准备： 

1、还原剂： 

3g NaBH4 +1g NaOH 定溶于100mL 的容量瓶中。

或 6g NaBH4 +2g NaOH 定溶于200mL 的容量瓶中。

或 9g NaBH4 +3g NaOH 定溶于300mL 的容量瓶中。

*还原剂每天需要配制
2、标准液：
 a)稀释液：，3%HNO3   空白液： 3%HNO3

 b)标准液, 如2ppb, 4ppb, 8ppb, 10ppb 等。
二、仪器操作规程： 

1、将氢化物石英池单元放在火焰燃烧头上，连接好氩气管路，并把传感器线与主机SFS连接口相连。把新配制的NaBH4溶液装好。
2、打开氩气瓶总阀，调节压力减压阀使气体出口压力为0.6MPa 。
3、打开计算机电源。
4、打开仪器电源，等仪器自检结束后，双击WINAAS 图标，选择Flame， Aplication/Cookbook ，载入汞分析程序，进入仪器操作系统。 如果己建立测试方法，可选择METHODE，
[image: image1.png]WinAAS Technique: Flame  Versio

317.0 analytikjena

Technique selection: Task selection:

e Date  [2007-06-05
Applications{cookbook Time 23:28

Methods

© Sequ. mult. rout

o

v

) © Development

Operator: \

Lab.: [

Code: [

Notes:

I~ Bimuiation

HoL turet Delete Reset Exit application oK





[image: image2.png]nt selection cookbook

I~ Emission

- Elements;
U | Be B c| w| ol F|ne
Na | Mg afsi| e s]o| a

Ro| s | v | zr | No| mo| | Ru|Rh| Pd| ag | Cd| In| sn| sb| Te *e

cs | Ba| taf | Tal w|Rre|os| | ptyu|lHg| ™) Po[ Bi| rof A en

Fr| Ra

Ce | Pr| Nd| Fr|sm| Eul od| To| Dy | Ho| Er| Tm| vb| Lu

Th| Pa| U | np| Pu| am|em| ek | cf | Es | Fm| md| o | Lr

Cancel





   [image: image3.png]Technique:  [/1ame (Abs] Element: Hg

—  Performance: Mercury

Charact. concentratian [mgfL1%Abs]
Check concentration (0.1 Abs ") [mofL]

Typioal valus (70%) or 100mm-burner and C2H2ja flame

— Spectrometer ——— Atomizer: Flame ———————————
Mainline  [nm] : 253.7 Flame CoH2AAIr
Afternat. line [ f/o"f“”‘ch‘”memc U o
Slitwidth [ © 05 uellow (L)

Usable burner height [mm] 5-10
HCL current  [ma] : 3.0

Alternative flame CoHINZO

< > Print Load Close





点“LOAD”装载方法.
5、在SPE CTROMETER 中调节汞空心阴极灯的电流、能量和增益。

注意：把燃烧头高度设为：Burner Height=15 (目的是让元素灯发出的光完全通过石英管。在调节能量时可先用一小纸片看光是否完全通过石英管，

Burner Height可在5-15之间选择，调节后看光电倍增管负高压（EHT），EHT越低越好。

[image: image4.png]Spectroeter

Optical parameters  Energy / Gain [integration parameters | Scan parameters | Scan |

Energy levels: Gain:
HeL: 51 D2-HCL: EHT: 248 \
— w

| Peak pick AGC

— a — Lamp alignment:
0 o 3 4 W oadj
@
Start Save
urner height:
Automat
Reset

o]




2）在“Intergration parameters”,按以下条件设置：

[image: image5.png]Recommended settings:

Tntegration type

Integral value

Smoothing Golay-Sawitzky with 25 points>
Integration time 45s (Te:25s)
AZ time 2s

"Autosampler

Tnjection switch

)
ff

Hydride parameters:

=5

Element Purge time Reaction time Conc. NaBH, Acid

As 40s 8s 1% 3% HCI

Bi 15s 5s 1% 3% HCI

e 0 5s 1% @
Sb

Se 40s 10s 1% 3% HCI

Te 30s 15s 1.5% 3% HCI





注：Purge time 指Prewash(S):冲洗时间，此参数在测定时会弹出菜单：
[image: image6.png]Nanual working mode

Use the next reaction beaker:

\l) Sample 1

Prewash [s]: Reaction [s]:




 

[image: image7.png]Spectrometer.

Optical parameters | Energy / Gain Integration parameters ‘SE,,, parameters | Scan |

Integration mode: Measuring times:

-

Peak smoothing:

-
Average: 20 [5 points
- Golay-Savitzky:

I Disable AZ drift check

Opt.meas.

Reset

0K





6．在“Autosample mode”,
Autosampler type  :     No autosampler
Injection switch  :     Off

[image: image8.png]Autosample mode | | | | | |

Autos Injection switch:
No autosampler) I

CRSET T ASE1S (g " Activate SFS §

© Ass2 © AS52S e

o o g

/4

=
- g

-
-

Sample volume:
e 00 & s
-

Reset

oK.





7．在“Options”:选择：Peak graphics(主要看出峰峰形)
[image: image9.png]Activation |Setiings |

Integrated values

W Alphanumerical values
W Bar graph values

R Conc.values in curve

Measurement peaks

¥ Peak graphics

I” Peak overlay

I~ Graphical peak overview

Report windows:

e

Test Screen

0K





8、在CALIBRATION 的MODE 中选择：， 

1）STANDARD CALIBRATION
[image: image10.png]Calibration mode |Canditions | Stati

ics | Conc. input| Table | Curve parameters | Recalibration |

Compution / calibration method:

 {Absorbance mode [no calibration)

Standard calibration

© Bracketing calibration

© Method of add

Reset

0K





2）标准点数：

[image: image11.png]Calibration mode Con

ons |Statstics | Conc. input | Table | Curve parameters | |

Standards: Getting values by:
Making standards: & Measurement
 Using premixed standards € Manual input
3

Mix conditions:
3

No. of standards [

T

I~ Continue after curve calculation

o





3）在STATISTICS 中选择SIGMA 统计，测量次数2-3 次

[image: image12.png]Calibration

Calibration made | Conditions Statistics |Conc. input | Table | Curve parameters | |
st Grubbs ou
c ot c oot

g mean by (i)

© Median Statistics © On (Without flagging)

Measurement cycles:

Runs/Std.:

Blind runs{Std.: L

Statistical output:

 Standard devi:

ion (SD)

¥ Relat. standard deviation [RSD)

Reset

0K





4）标准溶液浓度：

[image: image13.png]Calibration mode | Con

5 | Conc. input Table |Curve parameters | |

Standards:
Name |

PKA
SD

PkH
SD

[RSD g

L |

Clear Table Check Run sample

Delete

Parameters Fit curve

0K





按START 开始测量，并根据仪器提示加入标准液和3%HNO3，加入量如标准溶液：3ml ，3%HNO3：2ml。

[image: image14.png]# WinA&S Yer: 3.17.0 - Applications/cookbook

AAS ZEEnit 700 Tech: Hydride - [ Neth: As disc. HS6x] =
Calibration

Calibration made | Conditions | Statistics | Conc. input Table |curve parameters | As

=

Cal-Std1
Cal-Std2
Cal-Std3
Cal-Std4

Start conditions — Humber of protocol files:

. Clean syst
Saving measurement results:

€ No result saving

& {Start a new report file

 Append results to the existing file

Kl

I Overwrite last report ile

Clear Table

Delete line

StartiAbs





[image: image15.png]Calibration

| | | Tabe | | |

Standards: : PR ]
Name | State | Pos. [ SD_[RSD g

Nanual working mode

. Use the next reaction beaker:
Ll | 1 Cal-Zero - 0 mg/L

Cancel





设定好：Prewash:0
          Reation:5
点“OK”进行零标准测定。

按提示依次把标准曲线做好，相关系数最好R>0.995

[image: image16.png]- - - - - Peak area Hg
Linear 193.7 nm
» -
15 - Slope: 5.01822 Abs/mg/L
sc0: 0.00146 mg/L

Char.conc.:  0.00087 mg/L /1%A

0

Absorbance

o 1 2 3 4
Cone. [mgiL ]





注：每次测量后，用水反复冲洗反应池，至少两遍。

    还原剂NaBH4仪器自动加入反应池；

样品（或标准溶液）和3%HNO3溶液：手动加入反应池。

加入量如：标准溶液：3ml ；3% HNO3：2ml
9、标准曲线做好后便可进行样品测试。

10、样品测试完后，关闭氩气总阀，退出软件，软件退出后，关闭仪器主机，关闭计算机，。
